Minutes from ASP Laboratory Team Meeting

Date: January 15th, 2026

Present: Marcos Lemes (UM) and Zoe Walker (UC)

Absent: Brent Else (UC), Carolin Löscher (SDU), David Capelle (DFO), and Kitte Gerlich (AU), Christien Laber (UiT), and Egon Frandsen (AU)

Referee: Marcos

This is meeting has been taken online.

We have discussed below topics:

- Marcos (UM) gave an update on ASP General meeting past November 2025.

  - The interlaboratory exercise was presented and some issues were identified during the exercise, such as:

    - The need to adjust the ionic strength of titrant accordingly to type of sample (fresh water or seawater);

    - The use of borosilicate bottle instead soda lime flask for the analysis;

    - Allow more time for equilibration during transferring sample/standard.

- Marcos (UM) presented of well-received main findings from Preservation Study during the ASP General meeting:

  - The borosilicate exetainers need to be combusted before usage, due to it to be closer results to true value (CRM);

  - The use the UV-C radiation can replace HgCl2 poisoning for TA and DIC analyzes;

  - The data suggests that sample can be stored for 60 days without any chemical preservation in borosilicate exetainers;

  - Apparently, the amount of HgCl2 added to preserve samples can complex with dissolved organic matter, underestimating TA and DIC values in fresh water. According to Mos, B. et al. (2021). Alkalinity of diverse water samples can be altered by mercury preservation and borosilicate vial storage. Scientific Reports, 11(1). This underestimation can be seen in this study.

- Marcos (UM) and Zoe (UC) discussed a new schedule for the meeting and was decided to have meetings each two months. IF any member needs to have any discussition in this interval, may call for a meeting.

Please use Teams link https://teams.microsoft.com/l/meetup-join/19%3aiaBWS_EnwxKky61JDxNP89AT3vq9AlgyUTHa4FP4Xpc1%40thread.tacv2/1681337140562?context=%7b%22Tid%22%3a%224f80dd0b-338c-4e4c-8a14-90446962f7b8%22%2c%22Oid%22%3a%2271262080-0ca9-4c08-ba54-444a9d0cbce3%22%7d for the meetings

IMPORTANT NOTE: The next meeting will be March 12nd, 2026